- GENERAL

  - DESCRIPTION

    - Section Includes

- 1.	Analyzers, cables, and associated equipment for measuring free or total residual chlorine in water and wastewater treatment systems.

    - Scope

      - Furnish, install, calibrate, test, adjust, and place into satisfactory operation analyzers as shown on the Drawings and specified herein.

      - The Drawings and Specifications illustrate and specify functional and general construction requirements of analyzers and do not necessarily show or specify all components, wiring, piping, and accessories required to make a completely integrated system. Provide all components, piping, wiring, accessories, and labor required for a complete and integrated system.

    - Coordination:  Coordinate with other suppliers for installation of all items specified herein and required to ensure the complete and proper interfacing of all the components and systems.

  - PERFORMANCE REQUIREMENTS

    - Design Requirements

      - Display:

        - 4x20 backlit LCD

      - Instrument dimensions:

        - 16.8 x 14.8 x 7.8 inches (42.7 x 37.6 x 19.8 cm)

      - Mounting:

        - Wall mount

      - Instrument weight:

        - <18 lbs. (8 kg.)

      - Enclosure rating:

        - IP 66 (with door latched)

    - Performance Requirements

      - Measurement range:

        - 0 to 5 mg/L free or total residual chlorine

      - Accuracy:

        - ±5% or ±0.03 mg/L, whichever is greater

      - Precision:

        - ±3% or ±0.01 mg/L, whichever is greater

      - Limit of Detection:

        - 0.03 mg/L

      - Resolution

        - ±3% or ±0.01 mg/L, whichever is greater

      - Cycle time

        - Programmable; 2.5 to 60 minutes

    - Environmental Requirements

      - Operating Temperature:

        - 41 – 131 °F (5 – 55 °C)

      - Relative Humidity:

        - 90% at 40 °C (non-condensing)

    - Electrical Requirements:

      - Power:

        - 115-230 VAC , 50-60 Hz, 70VA

      - Relay(s):

        - Two 30 VDC, 6A; NO/NC; SPDT, Potential-free

      - Analog output(s):

        - 4-20 mA configurable with alarm switch

      - Digital output(s):

        - RS-485 Modbus RTU protocol or ASCII standard

  - QUALITY ASSURANCE

    - Acceptable Manufacturers:

      - Furnish DPD analyzer by the named manufacturers.

      - The named manufacturers have been specified to establish the standard of quality and performance of the equipment to be supplied.

      - Manufacturer shall be ISO 9001 certified.

  - PRODUCT DELIVERY, STORAGE, AND HANDLING

    - The chlorine measurement system shall not be delivered to the site until all product information and system shop drawings for the system have been approved.

    - Delivery and Acceptance Requirements:  Deliver materials to site in manufacturer’s original, unopened containers and packaging, with labels clearly identifying product name and manufacturer.

    - Storage and Handling Requirements:

      - Store and handle materials in accordance with manufacturer’s instructions.

      - Keep materials in manufacturer’s original, unopened containers and packaging until installation.

      - Store materials in clean, dry area indoors.

      - Protect materials during storage, handling, and installation to prevent damage.

        - Temperature range for storage: 41 – 158 °F (5 – 70 °C)

  - SUBMITTALS

    - Product data

      - Analyzer operating manual.

      - Mounting bracket installation instructions.

    - Manufacturer’s Certifications:  Submit manufacturer’s certification that materials comply with specified requirements and are suitable for intended application.

    - Warranty documentation:  Submit manufacturer’s standard warranties.

- PRODUCTS

  - GENERAL

    - Measurement principle:

      - The chlorine analyzer will employ N, N-Diethyl-p-phenylenediamine (DPD) method to measure free and total residual chlorine

    - Chlorine measurement system shall be intended for use as a stand-alone instrument or as part of a multi-parameter monitoring network by connecting chlorine analyzer to expansion modules specified in Section 2.3.F and network cable specified in Section 2.3.F.

  - MANUFACTURER

    - Provide products from the following manufacturer:

      - YSI Incorporated | 1700/1725 Brannum Lane | Yellow Springs, OH 45387 USA | +1-937-688-4255 | +1 877-726-0975 (US) | info@ysi.com

  - MANUFACTURED UNIT

    - Enclosure:

      - Single-cabin design

      - Single channel,

      - One measuring range

    - Flow cell Assembly:

      - Optics

        - Class 1 LED

        - Wavelength: 525 nm; adjustable 475 nm to 625 nm

        - Path length: > 1 cm

      - Static mixer

        - Using principle of Flow Injection Analysis

    - Sample pump assembly

    - Reagent pump assembly

    - Fan assembly

    - Terminal block

      - Fuse-.5A 250VAC 3AG SLO BLO

    - SNCIQ-x,x [SNCIQ-x,x SW] network connection cable.

      - IC2 module

      - Capable of transmitting digital communications and low voltage (24V) power.

      - Integrated shield.

Note to specifier:  Network cable is suitable to connect all components of the multiparameter measuring system.  Specify appropriate total length of cable required for project in meters.

  - MATERIALS

    - Enclosure

      - Polycarbonate

    - Flow cell Assembly

      - CPVC

    - Sample pump assembly

      - PA12 is Polyamide12

      - POM is Polyacetal.

    - Reagent pump assembly

      - Cyanoacrylate body

      - Stainless steel rollers

    - Fan assembly

      - Acrylonitrile Butadiene Styrene (ABS)

    - Terminal block

      - Polyamide 66 (PA 66).

  - ACCESSORIES

    - Sample inlet device

    - Wall-mount kit

- EXECUTION

  - PREPARATION

    - Mounting options

      - Benchtop rack

      - Wall mount

    - Enclosure space requirements

      - Vertical: 20.8 inches (52.9 cm)

      - Horizontal:18.8 inches (47.8 cm)

      - Depth: 11.4 inches (29 cm)

      - Door: 180° of clearance

    - Sample inlet

      - 1/8-inch (3.175-mm)

    - Sample outlet

      - ½-inch (12.7-mm) flexible tubing

  - INSTALLATION

    - Refer to the following documents:

      - Analyzer manual

      - Analyzer quick start guide

  - START-UP

    - Install analyzers in strict accordance with the manufacturer’s instructions and recommendations.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require representatives of the following be present during the start-up:

        - General contractor

        - Electrical contractor

        - Xylem factory-trained representative

        - Owner’s personnel

        - Engineer

- END OF SECTION