GENERAL

  - DESCRIPTION

    - Section Includes

      - Sensors, cables, and associated equipment for measuring conductivity in water and wastewater treatment systems.  Controllers and modules required for a complete monitoring system are specified elsewhere.

    - Scope

      - Furnish, install, calibrate, test, adjust, and place into satisfactory operation conductivity sensors as shown on the Drawings and specified herein.

      - The Drawings and Specifications illustrate and specify functional and general construction requirements of conductivity sensors and do not necessarily show or specify all components, wiring, piping, and accessories required to make a completely integrated system. Provide all components, piping, wiring, accessories, and labor required for a complete and integrated system.

    - Coordination:  Coordinate with other suppliers for installation of all items specified herein and required to ensure the complete and proper interfacing of all the components and systems.

  - PERFORMANCE REQUIREMENTS

    - Design Requirements

      - Design conductivity measurement system for continuous monitoring in situ with automatic temperature compensation that is adjustable by the user.

      - Conductivity measurement system shall be a component of a process monitoring system capable of monitoring up to 20 parameters simultaneously on each controller.  Measurement locations shall be networked together in a linear, star, or branched topology with a single 2-wire cable that distributes 12 V power and communications.

Note to specifier:  Item 1.2.A.2 is for a System 2020XT monitoring system.

      - The conductivity sensor assembly shall be a separate unit from the sensor connection cable and able to be located anywhere in the monitoring network a sensor connection cable is available and wired into the network.

    - Performance Requirements

      - Operating range

        - Temperature:  23°F to 140°F (-5°C to 60°C)

        - Pressure range:  up to 10 bar

        - pH:  4 S.U. to 12 S.U.

      - Measuring range

        - Conductivity:  10 µS / cm to 500 ms / cm

        - Salinity:  0.0 ppt to 70 ppt

        - Total dissolved solids (TDS):  0 mg/L to 2,000 mg / L

      - Accuracy

        - Conductivity, salinity, total dissolved solids: +/- 1.5%

        - Temperature:  +/- 0.5K

      - Resolution:

        - Conductivity

          - 0.00 to 20.00 μS/cm:  	0.01 μS / cm

          - 0.0 to 200.0 μS/cm:  	0.1 μS / cm

          - 0 to 2000 μS/cm: 		1 μS / cm

          - 0.00 to 20.00 mS/cm:  	0.01 mS / cm

          - 0.0 to 200.0 mS/cm:  	0.1 mS/cm

          - 0.0 to 500.0 mS/cm:  	0.1 mS/cm

        - Salinity: 0.1 ppt

        - Total dissolved solids: 1 mg/L

        - Temperature:  0.1 K

    - Response time

      - Temperature:

        - Response time t90: < 60 s

        - Response time t95:  < 120 s

  - QUALITY ASSURANCE

    - Acceptable Manufacturers:

      - Furnish conductivity sensors by the named manufacturers.

      - The named manufacturers have been specified to establish the standard of quality and performance of the equipment to be supplied.

      - Manufacturer shall be ISO 9001 certified.

  - PRODUCT DELIVERY, STORAGE, AND HANDLING

    - Conductivity measurement system shall not be delivered to the site until all product information and system shop drawings for the system have been approved.

    - Delivery and Acceptance Requirements:  Deliver materials to site in manufacturer’s original, unopened containers and packaging, with labels clearly identifying product name and manufacturer.

    - Storage and Handling Requirements:

      - Store and handle materials in accordance with manufacturer’s instructions.

      - Keep materials in manufacturer’s original, unopened containers and packaging until installation.

      - Store materials in clean, dry area indoors.

      - Protect materials during storage, handling, and installation to prevent damage.

      - Temperature range for storage:  23°F to 149°F (-5°C to 65°C)

  - SUBMITTALS

    - Product data

      - Sensor operating manual.

      - Mounting bracket / sensor holder installation instructions.

    - Manufacturer’s Certifications:  Submit manufacturer’s certification that materials comply with specified requirements and are suitable for intended application.

    - Warranty documentation:  Submit manufacturer’s standard warranties.

- PRODUCTS

  - GENERAL

    - Measurement principle shall be a 4-electrode measuring cell.

    - Connect conductivity sensor to field-mounted junction box or expansion modules specified in Section 40 95 00 with sensor connecting cable as specified in Section 2.3.A.2.

    - Conductivity measurement system components shall be designed to be part of an analytical process control system with enhanced protection from overvoltage due to lightning and power supply fluctuations according to EN 61326 when installed using manufacturer’s recommended components per manufacturer’s instructions.

  - MANUFACTURER

    - Provide products from the following manufacturer:

      - YSI, Incorporated, 1700/1725 Brannum Lane, Yellow Springs, OH  45387.  1-800-765-4974.

  - MANUFACTURED UNIT

    - The Conductivity measurement system consists of the following:

      - Model TetraCon® 700 IQ [Model TetraCon® 700 IQ SW] Conductivity Sensor.

        - The sensor shall use a four (4) electrode method that measures conductivity of the sample due to the presence of dissolved ions.

        - The sensor shall have a factory-calibrated cell constant.

          - K = 0.917 cm in immersion mount

          - K = 0.933 cm with flow-thru adapter defined in Section 2.5. of this specification.

        - The sensor shall have a digital output signal.

        - The sensor shall detach from sensor cable allowing for easy replacement or repair.

Note to specifier:  Select TetraCon® 700 IQ for measurements in water/wastewater applications.  Select TetraCon® 700 IQ SW for measurements in seawater/aquaculture and other highly corrosive environments.

      - SACIQ-x,x [SACIQ-x,x SW] sensor connection cable.

        - Capable of transmitting digital communications and low voltage (24V) power

        - Integrated shield.

        - Connections

          - Sensor-side:  Quick fastener (threaded).

          - Controller-side:  tinned conductors

Note to specifier:  Specify appropriate length of cable in meters, e.g. SACIQ-1,5 is 1.5 m.  Standard lengths are 1.5 m, 7m, and 15m.  Custom lengths up to 100 m are available.  Select SACIQ-x,x for measurements in water/wastewater applications.  Select SACIQ-x,x SW for measurements in seawater/aquaculture and other highly corrosive environments.

  - MATERIALS

    - Sensor

      - Body:  V4A Stainless Steel 1.4571 (equivalent to 316Ti)

Note to specifier:  Specify Titanium, POM for SW model applications.

      - Sensor head PVC, epoxy (filler)

      - Electrodes, housing of the temperature sensor: Graphite

      - Plug head connector housing:  POM

      - Plug, 3-pole:  ETFE (blue) Tefzel®

      - Safety

        - EN 61010-1

        - UL 3111-1

        - CAN/CSA C22.2 No. 1010.1

      - Warranty: 2 yrs.

    - Sensor connection cable

      - Conductors:  Tin-coated stranded copper wire.

      - Coupling ring: Stainless Steel 1.4571 (equivalent to 316Ti)

Note to specifier:  specify Titanium, grade 2 for SW model applications.

      - Ring: POM

Screw: Stainless Steel V4A

Note to specifier:  specify Polyamide 6.6 for SW model applications.

      - O-ring:  NBR

      - Enclosure:  POM

      - Nut:  Stainless Steel 1.4571 (equivalent to 316Ti)

Note to specifier:  specify Titanium, grade 2 for SW model applications.

      - Protection ring:  POM

      - Cable sheath:  PUR

      - Protective cap: PVC

Note to specifier:  specify POM threaded plug for SW model applications.

      - Protection rating:  IP68 (waterproof)

      - Warranty:  12 months

  - ACCESSORIES

Note to specifier:  Specify sensor mounting assembly and other accessories as required.

    - Handrail Sensor mounting assemblies:

      - YSI model 2458000Y SMK Quick Release Stainless Steel Handrail Mounting Bracket with 1.5-inch Mounting Adapter.

Note to specifier:  Requires 1.5-in. PVC pipe cut to desired length not included with mounting assembly from YSI.

    - Insertion sensor mounting assemblies

      - Model WA 700/x retractable armature

Note to specifier:  Specify WA 700/2 for 2 bar overpressure; WA/10 for 10 bar overpressure.

      - Model ADA-WA 1 sensor adapter

      - Welding socket

        - Model ESS-WA 700/xx

Note to specifier:  Specify ESS-WA 700/ST for St 33 steel; specify ESS-WA 700/VA for 1.4571 stainless steel.

- EXECUTION

  - INSTALLATION

    - Immersion sensors:

      - Immerse sensor a minimum of 4 in. (10 cm) below the water surface up to a maximum of 100 m below the water surface.

      - Protect the probe against the oncoming flow of large objects.

      - Install such that measuring end of the sensor is a minimum distance of 2 in. (5 cm) from floors and walls.

    - Insertion Sensors

      - Install in an up-flow pipe section for best results. Do not mount on the top or bottom of a horizontal pipe section.

      - Pipe diameter: minimum 2 inch (50 mm) diameter.

      - Install minimum 5 ft. (1.5 m) or three times the pipe diameter (whichever is greater) downstream of pumps, valves, or pipe elbows.

      - Install the sensor on the discharge side of a pump. To measure sludge with significant amounts of debris, install it after a sludge grinding pump.

    - Sensor connection cable

      - Bend radius

        - Permanent bend:  not less than 3.2 in. (80 mm).

        - One-time bend:  not less than 2 in. (50 mm).

  - START-UP

    - Install conductivity sensors in strict accordance with the manufacturer’s instructions and recommendations.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require representatives of the following be present during the start-up:

        - General contractor

        - Electrical contractor

        - YSI factory-trained representative

        - Owner’s personnel

        - Engineer

END OF SECTION