- GENERAL

  - DESCRIPTION

    - Section Includes

      - Sensors, cables, and associated equipment for measuring total suspended solids in wastewater treatment systems.  Controllers and modules required for a complete monitoring system are specified elsewhere.

    - Scope

      - Furnish, install, calibrate, test, adjust, and place into satisfactory operation total suspended solids measurement systems as shown on the Drawings and specified herein.

      - The Drawings and Specifications illustrate and specify functional and general construction requirements of total suspended solids measurement systems and do not necessarily show or specify all components, wiring, piping, and accessories required to make a completely integrated system. Provide all components, piping, wiring, accessories, and labor required for a complete and integrated system.

    - Coordination:  Coordinate with other suppliers for installation of all items specified herein and required to ensure the complete and proper interfacing of all the components and systems.

  - PERFORMANCE REQUIREMENTS

    - Design Requirements

      - Design total suspended solids measurement system for continuous monitoring in situ.

      - Total suspended solids measurement system shall be a component of a process monitoring system capable of monitoring up to 20 parameters simultaneously on each controller.  Measurement locations shall be networked together in a linear, star, or branched topology with a single 2-wire cable that distributes 12 V power and communications.

Note to specifier:  Item 1.2.A.2 is for a System 2020XT monitoring system.

      - The total suspended solids sensor assembly shall be a separate unit from the sensor connection cable and able to be located anywhere in the monitoring network a SensorNet terminal is available.

      - Total suspended solids measurement system(s), when used together with approved associated nonincendive field wiring apparatus, shall be designed for use in hazardous classified areas where explosive gases are not present under normal operating conditions.

Note to specifier:  Include Item 1.2.A.4 for applications in Class I Division 2 classified areas requiring Model ViSolid® 700 IQ H

    - Performance Requirements

      - Operating range

        - Temperature:  32°F to 140°F (0°C to 60°C)

        - Pressure range:  0.5 bar to 10 bar

        - pH:  4 S.U. to 12 S.U.

      - Measuring range:

        - 3 to 40,000 mg TSS / L.

        - 0.0003 to 100% TSS

      - Process variation coefficient

        - Low concentration: < 2%

        - High concentration: < 4%

      - Resolution:

        - 0 to 400.0 mg TSS / L ; 0 to 400.0 ppm TSS : 0.1 mg TSS / L ; 0.1 ppm

        - 0 to 4,000 mg TSS / L; 0 to 4,000 ppm TSS : 1 mg TSS / L ; 1 ppm

        - 0 to 25,000 mg TSS / L; 0 to 2.500 % TSS : 1 mg TSS / L ; 0.001 %

        - 0 to 40,000 mg TSS / L; 0 to 4.000 % TSS : 1 mg TSS / L ; 0.001 %

        - 0 to 40 % TSS :  0.01 %

        - 0 to 100.0 % TSS : 0.1 %

      - Signal averaging:  User selectable to provide a t90 from 1 to 600 seconds.

  - QUALITY ASSURANCE

    - Acceptable Manufacturers:

      - Furnish total suspended solids sensors by the named manufacturers.

      - The named manufacturers have been specified to establish the standard of quality and performance of the equipment to be supplied.

      - Manufacturer shall be ISO 9001 certified.

  - PRODUCT DELIVERY, STORAGE, AND HANDLING

    - Total suspended solids measurement system shall not be delivered to the site until all product information and system shop drawings for the system have been approved.

    - Delivery and Acceptance Requirements:  Deliver materials to site in manufacturer’s original, unopened containers and packaging, with labels clearly identifying product name and manufacturer.

    - Storage and Handling Requirements:

      - Store and handle materials in accordance with manufacturer’s instructions.

      - Keep materials in manufacturer’s original, unopened containers and packaging until installation.

      - Store materials in clean, dry area indoors.

      - Protect materials during storage, handling, and installation to prevent damage.

      - Temperature range for storage:  23°F to 149°F (-5°C to 65°C)

  - SUBMITTALS

    - Product data

      - Sensor operating manual.

      - Mounting bracket / sensor holder installation instructions.

      - Control drawing.

Note to specifier:  Include Item 1.5.A.3 for applications in Class I Division 2 classified areas requiring Model ViSolid® 701 IQ H

    - Manufacturer’s Certifications:  Submit manufacturer’s certification that materials comply with specified requirements and are suitable for intended application.

    - Warranty documentation:  Submit manufacturer’s standard warranties.

- PRODUCTS

  - GENERAL

    - Measurement principle shall be optical based on the reflectance of light from a near infrared light source.

    - Connect total suspended solids sensor to field-mounted junction box or expansion modules specified in Section 40 95 00 with sensor connecting cable as specified in Section 2.3.A.2.

    - Total suspended solids measurement system components shall be designed to be part of an analytical process control system with enhanced protection from overvoltage due to lightning and power supply fluctuations according to EN 61326 when installed using manufacturer’s recommended components per manufacturer’s instructions.

  - MANUFACTURER

    - Provide products from the following manufacturer:

      - YSI, Incorporated, 1700/1725 Brannum Lane, Yellow Springs, OH  45387.  1-800-765-4974.

  - MANUFACTURED UNIT

    - The Total suspended solids measurement system consists of the following:

      - Model ViSolid® 700 IQ [Model ViSolid® 700 IQ SW; Model ViSolid® 700 IQ H] Total Suspended Solids Sensor.

        - The sensor shall include detectors for measuring the intensity of scattered light and backscatter.

        - The sensor shall be factory-calibrated for low concentration (mixed liquor, return activated sludge) and high concentration (primary sludge, thickened sludge) applications.

        - The sensor shall have capability for automatically selecting the optimal resolution for each measured value.

        - The sensor shall include self-diagnostics to detect contamination of optical window and failure of cleaning system.

        - The sensor shall have a non-mechanical integrated ultrasonic cleaning system.

        - The sensor shall have a digital output signal.

        - The sensor shall detach from sensor cable allowing for easy replacement or repair.

        - The sensor shall be designed to be nonincendive and nonarcing.

Note to specifier:  Select ViSolid® 700 IQ for measurements in water/wastewater applications.  Select ViSolid® 700 IQ SW for measurements in seawater/aquaculture and other highly corrosive environments.  Select ViSolid® 700 IQ H for Class I Division 2 classified areas and include Item 2.2.A.1.h.

      - SACIQ-x,x [SACIQ-x,x SW] sensor connection cable.

        - Capable of transmitting digital communications and low voltage (24V) power

        - Integrated shield.

        - Connections

          - Sensor-side:  Quick fastener (threaded)

          - Controller-side:  tinned conductors

Note to specifier:  Specify appropriate length of cable in meters, e.g. SACIQ-1,5 is 1.5 m.  Standard lengths are 1.5 m, 7m, and 15m.  Custom lengths up to 100 m are available.  Select SACIQ-x,x for measurements in water/wastewater applications.  Select SACIQ-x,x SW for measurements in seawater/aquaculture and other highly corrosive environments.

  - MATERIALS

    - Sensor

      - Body:  V4A Stainless Steel 1.4571 (equivalent to 316Ti)

Note to specifier:  Specify Titanium, POM for SW model applications.

      - Measuring window:  sapphire

      - Certifications

        - Safety: CE,  cETLus

        - Hazardous locations:  CSA C22.2 No. 213, ANSI/ISA 12.12.01

Note to specifier:  Include Item 2.3.A.8.b for applications in Class I Division 2 classified areas requiring Model ViSolid® 700 IQ H

      - Protection rating:  IP68

      - Warranty: 2 yrs.

    - Sensor connection cable

      - Conductors:  Tin-coated stranded copper wire.

      - Coupling ring: Stainless Steel 1.4571 (equivalent to 316Ti)

Note to specifier:  specify Titanium, grade 2 for SW model applications.

      - Ring: POM

- Screw: Stainless Steel V4A

Note to specifier:  specify Polyamide 6.6 for SW model applications.

      - O-ring:  NBR

      - Enclosure:  POM

      - Nut:  Stainless Steel 1.4571 (equivalent to 316Ti)

Note to specifier:  specify Titanium, grade 2 for SW model applications.

      - Protection ring:  POM

      - Cable sheath:  PUR

      - Protective cap: PVC

Note to specifier:  specify POM threaded plug for SW model applications.

      - Protection rating:  IP68 (waterproof)

      - Warranty:  12 months

  - ACCESSORIES

Note to specifier:  Specify sensor mounting assembly and other accessories as required.

    - Handrail Sensor mounting assemblies:

      - YSI model 2458000Y SMK Quick Release Stainless Steel Handrail Mounting Bracket with 1.5-inch Mounting Adapter.

Note to specifier:  Requires 1.5-in. PVC pipe cut to desired length not included with mounting assembly from YSI.

    - YSI model 207002Y MIQ/BB2 barrier box

      - An Associated Nonincendive Field Wiring Apparatus shall be inserted between each total suspended solids measurement system and the monitoring system network.

Note to specifier:  Include Item 2.4.D for applications in Class I Division 2 classified areas requiring Model ViSolid® 700 IQ H

    - Insertion sensor mounting assemblies

      - Model WA 700/x retractable armature

Note to specifier:  Specify WA 700/2 for 2 bar overpressure; WA/10 for 10 bar overpressure.

      - Model ADA-WA 1 sensor adapter

      - Welding socket

        - Model ESS-WA 700/xx

Note to specifier:  Specify ESS-WA 700/ST for St 33 steel; specify ESS-WA 700/VA for 1.4571 stainless steel.

- EXECUTION

  - INSTALLATION

    - Immersion sensors:

      - Immerse sensor a minimum of 4 in. (10 cm) below the surface.

      - Install the sensor at an angle from vertical of 20° to 45° facing upstream.

Note to specifier:  the sensor can be installed facing downstream if there is a high incidence of fibrous or flat items, e.g. hair or leaves.

      - Protect the probe against the oncoming flow of large objects.

      - Install where the optical window does not face into direct light.

      - Install sensor so that window is a minimum distance of 4 in. (10 cm) from floors and walls.

    - Insertion Sensors

      - Install in an up-flow pipe section for best results. Do not mount on the top or bottom of a horizontal pipe section.

      - Pipe diameter: minimum 2 inch (50 mm) diameter.

      - Install minimum 5 ft. (1.5 m) or three times the pipe diameter (whichever is greater) downstream of pumps, valves, or pipe elbows.

      - Install the sensor on the discharge side of a pump. To measure sludge with significant amounts of debris, install it after a sludge grinding pump.

    - Sensor connection cable

      - Bend radius

        - Permanent bend:  not less than 3.2 in. (80 mm).

        - One-time bend:  not less than 2 in. (50 mm).

    - Install barrier box MIQ/BB2 outside the classified area.

    - Connect each hazardous location total suspended solids measurement system to a separate barrier box MIQ/BB2.

Note to specifier:  Include Items 3.1.D and 3.1.E for applications in classified areas requiring Model ViSolid® 700 IQ H.

  - START-UP

    - Install total suspended solids sensors in strict accordance with the manufacturer’s instructions and recommendations.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require representatives of the following be present during the start-up:

        - General contractor

        - Electrical contractor

        - YSI factory-trained representative

        - Owner’s personnel

        - Engineer

- END OF SECTION