- GENERAL

  - DESCRIPTION

    - Section Includes

      - Sensors, cables, and associated equipment for measuring turbidity in water and wastewater treatment systems.  Controllers and modules required for a complete monitoring system are specified elsewhere.

    - Scope

      - Furnish, install, calibrate, test, adjust, and place into satisfactory operation turbidity sensors as shown on the Drawings and specified herein.

      - The Drawings and Specifications illustrate and specify functional and general construction requirements of turbidity sensors and do not necessarily show or specify all components, wiring, piping, and accessories required to make a completely integrated system. Provide all components, piping, wiring, accessories, and labor required for a complete and integrated system.

    - Coordination:  Coordinate with other suppliers for installation of all items specified herein and required to ensure the complete and proper interfacing of all the components and systems.

  - PERFORMANCE REQUIREMENTS

    - Design Requirements

      - Design turbidity measurement system for continuous monitoring in situ.

      - Turbidity measurement system shall be a component of a process monitoring system capable of monitoring up to 20 parameters simultaneously on each controller.  Measurement locations shall be networked together in a linear, star, or branched topology with a single 2-wire cable that distributes 12 V power and communications.

Note to specifier:  Item 1.2.A.2 is for a System 2020XT monitoring system.

      - The turbidity sensor assembly shall be a separate unit from the sensor connection cable and able to be located anywhere in the monitoring network a SensorNet terminal is available.

    - Performance Requirements

      - Operating range

        - Temperature:  32°F to 140°F (0°C to 60°C)

        - Pressure range:  0.5 bar to 10 bar

        - pH:  4 S.U. to 12 S.U.

      - Measuring range: 0.05 to 4000 FNU/NTU/TEF; 0.0001 to 400 g TSS / l

      - Process variation coefficient:  < 1% up to 2000 FNU

      - Resolution:

        - 0 to 400.0 mg TSS / l ; 0 to 400.0 ppm TSS : 0.1 mg TSS / l ; 0.1 ppm

        - 0 to 4,000 mg TSS / l ; 0 to 4,000 ppm TSS : 1 mg TSS / l ; 1 ppm

        - 0 to 25.00 g TSS / l ; 0 to 2.500 % TSS : 0.01 g TSS / l ; .001 %

        - 0 to 40.00 g TSS / l ; 0 to 4.000 % TSS : 0.01 g TSS / l ; .001 %

        - 0 to 400.00 g TSS / l ; 0 to 4.000 % TSS : 0.1 g TSS / l ; .01 %

        - 0 to 1,000 g TSS / l ; 0 to 100.0 % TSS : 1 g TSS / l ; .1 %

      - Signal averaging:  User selectable to provide a t90 from 1 to 600 seconds.

  - QUALITY ASSURANCE

    - Acceptable Manufacturers:

      - Furnish turbidity sensors by the named manufacturers.

      - The named manufacturers have been specified to establish the standard of quality and performance of the equipment to be supplied.

      - Manufacturer shall be ISO 9001 certified.

  - PRODUCT DELIVERY, STORAGE, AND HANDLING

    - Delivery and Acceptance Requirements:  Deliver materials to site in manufacturer’s original, unopened containers and packaging, with labels clearly identifying product name and manufacturer.

    - Storage and Handling Requirements:

      - Store and handle materials in accordance with manufacturer’s instructions.

      - Keep materials in manufacturer’s original, unopened containers and packaging until installation.

      - Store materials in clean, dry area indoors.

      - Protect materials during storage, handling, and installation to prevent damage.

      - Temperature range for storage:  23°F to 149°F (-5°C to 65°C)

  - SUBMITTALS

    - Product data

      - Sensor operating manual.

      - Mounting bracket / sensor holder installation instructions.

    - Manufacturer’s Certifications:  Submit manufacturer’s certification that materials comply with specified requirements and are suitable for intended application.

    - Warranty documentation:  Submit manufacturer’s standard warranties.

- PRODUCTS

  - GENERAL

    - Measurement principle shall be nephelometry

    - Connect turbidity sensor to field-mounted expansion modules specified in Section 40 95 00 with sensor connecting cable as specified in Section 2.3.A.2.

    - Turbidity measurement system components shall be designed to be part of an analytical process control system with enhanced protection from overvoltage due to lightning and power supply fluctuations according to EN 61326 when installed using manufacturer’s recommended components per manufacturer’s instructions.

  - MANUFACTURER

    - Provide products from the following manufacturer:

      - YSI, Incorporated, 1700/1725 Brannum Lane, Yellow Springs, OH  45387.  1-800-765-4974.

  - MANUFACTURED UNIT

    - The Turbidity measurement system consists of the following:

      - Model VisoTurb® 700 IQ [Model VisoTurb® 700 IQ SW] Turbidity Sensor.

        - The sensor shall use an optical method that measures the intensity of scattered light and backscatter from an LED source.

        - The sensor shall be factory-calibrated for turbidity.

        - The sensor shall have capability for automatically selecting the optimal resolution for each measured value.

        - The sensor shall include self-diagnostics to detect contamination of optical window and failure of cleaning system.

        - The sensor shall have a non-mechanical integrated ultrasonic cleaning system.

        - The sensor shall have a digital output signal.

        - The sensor shall detach from sensor cable allowing for easy replacement or repair.

Note to specifier:  Select VisoTurb® 700 IQ for measurements in water/wastewater applications.  Select VisoTurb® 700 IQ SW for measurements in seawater/aquaculture and other highly corrosive environments.

      - SACIQ-x,x [SACIQ-x,x SW] sensor connection cable.

        - Capable of transmitting digital communications and low voltage (24V) power

        - Integrated shield.

        - Connections

          - Sensor-side:  Quick fastener (threaded).

          - Controller-side:  tinned conductors

Note to specifier:  Specify appropriate length of cable in meters, e.g. SACIQ-1,5 is 1.5 m.  Standard lengths are 1.5 m, 7m, and 15m.  Custom lengths up to 100 m are available.  Select SACIQ-x,x for measurements in water/wastewater applications.  Select SACIQ-x,x SW for measurements in seawater/aquaculture and other highly corrosive environments.

  - MATERIALS

    - Sensor

      - Body:  V4A Stainless Steel 1.4571 (equivalent to 316Ti)

Note to specifier:  Specify Titanium, POM for SW model applications.

      - Measuring window:  sapphire

      - Safety

        - Certification: CE,  cETLus

      - Warranty: 2 yrs.

    - Sensor connection cable

      - Conductors:  Tin-coated stranded copper wire.

      - Coupling ring: Stainless Steel 1.4571 (equivalent to 316Ti)

Note to specifier:  specify Titanium, grade 2 for SW model applications.

      - Ring: POM

- Screw: Stainless Steel V4A

Note to specifier:  specify Polyamide 6.6 for SW model applications.

      - O-ring:  NBR

      - Enclosure:  POM

      - Nut:  Stainless Steel 1.4571 (equivalent to 316Ti)

Note to specifier:  specify Titanium, grade 2 for SW model applications.

      - Protection ring:  POM

      - Cable sheath:  PUR

      - Protective cap: PVC

Note to specifier:  specify POM threaded plug for SW model applications.

      - Protection rating:  IP68 (waterproof)

      - Warranty:  12 months

  - ACCESSORIES

Note to specifier:  Specify sensor mounting assembly and other accessories as required.

    - Handrail Sensor mounting assemblies:

      - YSI model 2458000Y SMK Quick Release Stainless Steel Handrail Mounting Bracket with 1.5-inch Mounting Adapter.

Note to specifier:  Requires 1.5-in. PVC pipe cut to desired length not included with mounting assembly from YSI.

    - Insertion sensor mounting assemblies

      - Model WA 700/x retractable armature

Note to specifier:  Specify WA 700/2 for 2 bar overpressure; WA/10 for 10 bar overpressure.

      - Model ADA-WA 1 sensor adapter

      - Welding socket

        - Model ESS-WA 700/xx

Note to specifier:  Specify ESS-WA 700/ST for St 33 steel; specify ESS-WA 700/VA for 1.4571 stainless steel.

- EXECUTION

  - INSTALLATION

    - Immersion sensors:

      - Immerse sensor a minimum of 4 in. (10 cm) below the surface.

      - Immerse sensor a maximum of 100 m below the surface.

      - Install the sensor at an angle from vertical of 20° to 45° facing upstream.

Note to specifier:  the sensor can be installed facing downstream if there is a high proportion of fibrous or flat items, e.g. hair or leaves.

      - Protect the probe against the oncoming flow of large objects.

      - Install where the optical window does not face into direct light.

      - Install sensor so that window is a minimum distance of 4 in. (10 cm) from floors and walls.

    - Insertion Sensors

      - Install at an angle of 45° against the direction of flow.

      - Install in an up-flow pipe section for best results. Do not mount on the top or bottom of a horizontal pipe section.

      - Pipe diameter: minimum 2 inch (50 mm) diameter.

      - Install minimum 5 ft. (1.5 m) or three times the pipe diameter (whichever is greater) downstream of pumps, valves, or pipe elbows.

      - Ensure a minimum of 10 in. (25 cm) of straight pipe upstream from sensor.

    - Sensor connection cable

      - Bend radius

        - Permanent bend:  not less than 3.2 in. (80 mm).

        - One-time bend:  not less than 2 in. (50 mm).

  - START-UP

    - Install turbidity sensors in strict accordance with the manufacturer’s instructions and recommendations.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require representatives of the following be present during the start-up:

        - General contractor

        - Electrical contractor

        - YSI factory-trained representative

        - Owner’s personnel

        - Engineer

- END OF SECTION